## Meter_Calibration
cal_row | instrument | calibration_date | as_found_factor | as_left_factor | units | linearity_error_pct | certificate_id | authority
CAL-001 | FT-CW241B | 2026-06-16 | 0.977 | 0.973 | dimensionless | 0.42 | CAL-CW241B-2026-06 | ACCEPTED
CAL-002 | FT-CW241B | 2024-03-08 | 1 | 1 | dimensionless | 0.95 | CAL-CW241B-2024-03 | SUPERSEDED
CAL-003 | Portable ultrasonic clamp-on | 2026-06-18 |  | 0.989 | dimensionless | 1.1 | PU-1187-CHECK | CROSSCHECK

## Trend
timestamp | raw_flow_gpm | current_factor | corrected_flow_gpm | register_factor | legacy_corrected_gpm | delta_gpm | basis_note
2026-06-18 00:00 | 1415 | 0.973 | 1376.8 | 1 | 1415 | -38.2 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 01:00 | 1419 | 0.973 | 1380.7 | 1 | 1419 | -38.3 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 02:00 | 1423 | 0.973 | 1384.6 | 1 | 1423 | -38.4 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 03:00 | 1427 | 0.973 | 1388.5 | 1 | 1427 | -38.5 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 04:00 | 1431 | 0.973 | 1392.4 | 1 | 1431 | -38.6 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 05:00 | 1435 | 0.973 | 1396.3 | 1 | 1435 | -38.7 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 06:00 | 1415 | 0.973 | 1376.8 | 1 | 1415 | -38.2 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 07:00 | 1419 | 0.973 | 1380.7 | 1 | 1419 | -38.3 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 08:00 | 1423 | 0.973 | 1384.6 | 1 | 1423 | -38.4 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 09:00 | 1427 | 0.973 | 1388.5 | 1 | 1427 | -38.5 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 10:00 | 1431 | 0.973 | 1392.4 | 1 | 1431 | -38.6 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 11:00 | 1435 | 0.973 | 1396.3 | 1 | 1435 | -38.7 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 12:00 | 1415 | 0.973 | 1376.8 | 1 | 1415 | -38.2 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 13:00 | 1422 | 0.973 | 1383.6 | 1 | 1422 | -38.4 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 14:00 | 1426 | 0.973 | 1387.5 | 1 | 1426 | -38.5 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 15:00 | 1430 | 0.973 | 1391.4 | 1 | 1430 | -38.6 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 16:00 | 1434 | 0.973 | 1395.3 | 1 | 1434 | -38.7 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 17:00 | 1438 | 0.973 | 1399.2 | 1 | 1438 | -38.8 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 18:00 | 1418 | 0.973 | 1379.7 | 1 | 1418 | -38.3 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 19:00 | 1422 | 0.973 | 1383.6 | 1 | 1422 | -38.4 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 20:00 | 1426 | 0.973 | 1387.5 | 1 | 1426 | -38.5 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 21:00 | 1430 | 0.973 | 1391.4 | 1 | 1430 | -38.6 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 22:00 | 1434 | 0.973 | 1395.3 | 1 | 1434 | -38.7 | Current factor from calibration certificate CAL-CW241B-2026-06.
2026-06-18 23:00 | 1438 | 0.973 | 1399.2 | 1 | 1438 | -38.8 | Current factor from calibration certificate CAL-CW241B-2026-06.

## Meter_Config
config_item | current_value | units | source | status | note
line_size | 10 | in | as-built spool tag | ACCEPTED | Same size as meter tube.
fluid | treated cooling water | text | process spec CW-04 | ACCEPTED | Specific gravity near 1.00.
transmitter_range | 2500 | gpm | instrument loop sheet | ACCEPTED | Used for raw indication.
current_factor | 0.973 | factor | CAL-CW241B-2026-06 | ACCEPTED | Applied to raw shift log flow.
register_factor | 1 | factor | asset register Rev 10 | SUPERSEDED | Do not use after 2026 calibration.

## Flow_Reconciliation
item | raw_value | factor | calculated_value | units | comparison_value | variance | status | source_trace
steady raw flow | 1428 | 0.973 | 1389.4 | gpm | 1450 | -4.2 | LOW_AGAINST_DESIGN | shift log + calibration workbook
highest corrected flow | 1441 | 0.973 | 1402.1 | gpm | 1450 | -3.3 | LOW_AGAINST_DESIGN | trend row 2026-06-18 23:00
legacy factor example | 1428 | 1 | 1428 | gpm | 1450 | -1.5 | SUPERSEDED_METHOD | register factor is superseded